## Leachate
KCSWD - Laboratory Leachate Data Report QA Audit
Site: |  | Cedar Hills |  |  |  | Monitoring Wells/Stations: |  |  |  | RIFS
Report Date: |  | 2015-07-14 |  |  |  | COC (SDG) #: |  |  |  | 13275
Sample Date(s): |  | 2015-06-16 |  |  |  | Lab Sample #: |  |  |  | L62924-2
Reviewed By: |  | ks |  |  |  | Field QA/QC Samples:
Hold Time Analysis Completed?: |  | Yes |  |  |  | (Avg) Sample Cooler Temp(s) (°C): |  |  |  | 3
Historical Data Comparison Completed?: |  | Yes |  |  |  | # Analytes Validated; Delivered: |  |  |  | 80 | 80
Lab/Field Data Comparison Completed?: |  | N/A |  |  |  | Data Completeness (%): |  |  |  | 100%
Summary of KCEL Laboratory Audit Report (note any anomalies reported from the cover page(s) of lab report):
A field-filtered container for dissolved metals analysis was not collected for this sample, therefore only total metals analysis was conducted. a low concentration of chloromethane was detected in the method blank (MB). Chloromethane results for the MB and sample LFS3150616- were qualified B. Also, recoveries of bromomethane from the spike blank (SB) and chloromethane from the matrix spike (MS) and MS duplicate (MSD) were below the lower control limit. The MS/MSD were prepared using a SWD sample collected from another location.
Conventionals | Storage / Preservation Techniques | Prep / Analysis Hold Times | Analyzed within Hold Time? (Y/N) | Method | Reporting Limits1 | Method Blank | Spike Blank (% Rec) | SB (% Rec Limits) | Matrix Spike (% Rec) | MS (% Rec Limits) | Lab Duplicates (% Rec) | LD (RPD % Limit) | Lab Control Samples (% Rec) | LCS (% Rec Limits) | Check Standard | CS (% Rec Limits) | QAPP Compliant?
Alkalinity, Bicarbonate | NA / Cool to ≤6°C w/in 15 min | NA / 14 Days | Y | SM 2320B | 1 mg/L as CaCO3 | [5 empty] | Y | 0-10% |  |  |  |  | Y
Alkalinity | NA / Cool to ≤6°C w/in 15 min | NA / 14 Days | Y | SM 2320B | 1 mg/L as CaCO3 | [5 empty] | Y | 0-10% | Y | 85-115% (>=50) or 90-110% (<=25) |  |  | Y
Ammonia Nitrogen | Cool to ≤6°C / Filter & Freeze to -20° C w/in 1 Day | NA / 28 Days | Y | SM 4500-NH3-G | 0.01 mg/L | Y | Y | 80-120% | Y | 75-125% | Y | 0-20% | Y | 85-115% |  |  | Y
Nitrate Nitrogen | Cool to ≤6°C / Filter & Cool to 6° C w/in 1 Day | NA / 2 Days | Y | SM 4500-NO3-F | 0.01 mg/L | Y | Y | 80-120% | Y | 75-125% | Y | 0-20% | Y | 85-115% |  |  | Y
Nitrite +Nitrate Nitrogen | Cool to ≤6°C / Filter & Cool to 6° C w/in 1 Day | NA / 28 Days | Y* | SM 4500-NO23-F | 0.01 mg/L | Y | Y | 80-120% | Y | 75-125% | Y | 0-20% | Y | 85-115% |  |  | Y
Chloride | Cool to ≤6°C | NA / 28 Days | Y | SM 4110B CL | 0.1 mg/L | Y | Y | 80-120% | Y | 75-125% | Y | 0-20% | Y | 85-115% | Y | 70-130% | Y
Sulfate | NA / Cool to ≤6°C w/in 15 min | NA / 28 Days | Y | SM 4110B SO4 | 0.1 mg/L | Y | Y | 80-120% | Y | 75-125% | Y | 0-20% | Y | 85-115% | Y | 70-130% | Y
Total Organic Carbon | Add H3PO4 to pH < 2 Immediately / Cool to ≤6°C | NA / 28 Days | Y | SM 5310B | 1 mg/L | Y | Y | 80-120% | Y | 75-125% | Y | 0-20% | Y | 85-115% |  |  | Y
Total Dissolved Solids | NA / Cool to ≤6°C w/in 15 min | NA / 7 Days | Y | SM 2540C | 5 mg/L | Y |  |  |  |  | Y | 0-25% | Y | 80-120% |  |  | Y
Total Suspended Solids | NA / Cool to ≤6°C w/in 15 min | NA / 7 Days | Y | SM 2540D | 1 mg/L | Y |  |  |  |  | Y | 0-25% | Y | 80-120% |  |  | Y
Total Metals | Storage / Preservation Techniques | Prep / Analysis Hold Times | Analyzed within Hold Time? (Y/N) | Method | Reporting Limits1 | Method Blank | Spike Blank (% Rec) | SB (% Rec Limits) | Matrix Spike (% Rec) | MS (% Rec Limits) | Lab Duplicates (% Rec) | LD (RPD % Limit) |  |  |  |  | QAPP Compliant?
Aluminum, Tot, ICP | Add HNO3 to pH <2. Transport on ice. | NA / 6 months | Y | SW 6010C | 0.01 mg/L | Y | Y | 85-115% | Y | 75-125% | Y | 0-20% |  |  |  |  | Y
Antimony, Tot, ICP |  |  |  |  | 0.01 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Arsenic, Tot, ICP |  |  |  |  | 0.001 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Barium, Tot, ICP |  |  |  |  | 0.015 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Beryllium, Tot, ICP |  |  |  |  | 0.001 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Cadmium, Tot, ICP |  |  |  |  | 0.01 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Calcium, Tot, ICP |  |  |  |  | 0.3 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Chromium, Tot, ICP |  |  |  |  | 0.003 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Cobalt, Tot, ICP |  |  |  |  | 0.05 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Copper, Tot, ICP |  |  |  |  | 0.001 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Hardness, ICP |  |  |  |  | 0.01 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Iron, Tot, ICP |  |  |  |  | 0.001 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Lead, Tot, ICP |  |  |  |  | 0.001 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Magnesium, Tot, ICP |  |  |  |  | 0.001 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Manganese, Tot, ICP |  |  |  |  | 0.001 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Nickel, Tot, ICP |  |  |  |  | 0.002 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Potassium, Tot, ICP |  |  |  |  | 0.005 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Selenium, Tot, ICP |  |  |  |  | 0.003 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Silver, Tot, ICP |  |  |  |  | 0.002 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Sodium, Tot, ICP |  |  |  |  | 0.004 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Thallium, Tot, ICP |  |  |  |  | 0.004 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Tin, Tot, ICP |  |  |  |  | 0.004 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Vanadium, Tot, ICP |  |  |  |  | 0.01 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Zinc, Tot, ICP |  |  |  |  | 0.001 mg/L |  |  | 85-115% |  | 75-125% |  | 0-20%
Mercury, Tot, CVAA | Add HNO3 to pH <2. Transport on ice. | 28 Days/ 28 Days | Y | SW 7470A | 0.0001 mg/L | Y | Y | 85-115% | Y | 75-125% | [6 empty] | Y
Volatile Organics | Storage / Preservation Techniques | Prep / Analysis Hold Times | Analyzed within Hold Time? (Y/N) | Method | Reporting Limits1 | Method Blank | Spike Blank (% Rec) | SB (% Rec Limits) | Matrix Spike & Matrix Spike Duplicate (% Rec) | MS & MSD (% Rec Limits) | MS/MSD (% Rec) | MS/MSD (RPD % Limit) |  |  |  |  | QAPP Compliant?
1,1,1,2-Tetrachloroethane | NA / Cool to ≤6°C w/in 15 min, can Add HCl to Increase Holding Time from 7 to 14 Days | NA / 14 days | Y | SW8260C | 0.2 ug/L | Y | Y | 80-125% | Y | 80-127% | Y | 0-40% |  |  |  |  | Y
1,1,1-Trichloroethane |  |  |  |  | 0.2 ug/L |  |  | 80-129% |  | 80-123% |  | 0-40%
1,1,2,2-Tetrachloroethane |  |  |  |  | 0.2 ug/L |  |  | 80-129% |  | 80-132% |  | 0-40%
1,1,2-Trichloroethane |  |  |  |  | 0.2 ug/L |  |  | 80-122% |  | 80-123% |  | 0-40%
1,1-Dichloroethane |  |  |  |  | 0.2 ug/L |  |  | 80-127% |  | 80-124% |  | 0-40%
1,1-Dichloroethene |  |  |  |  | 0.2 ug/L |  |  | 80-135% |  | 80-127% |  | 0-40%
1,2,3-Trichloropropane |  |  |  |  | 0.2 ug/L |  |  | 80-121% |  | 80-122% |  | 0-40%
1,2-Dibromo-3-Chloropropane |  |  |  |  | 1 ug/L |  |  | 76-122% |  | 76-128% |  | 0-40%
1,2-Dibromoethane |  |  |  |  | 0.2 ug/L |  |  | 80-123% |  | 80-124% |  | 0-40%
1,2-Dichlorobenzene |  |  |  |  | 0.2 ug/L |  |  | 80-129% |  | 80-128% |  | 0-40%
1,2-Dichloroethane |  |  |  |  | 0.2 ug/L |  |  | 80-128% |  | 80-127% |  | 0-40%
1,2-Dichloropropane |  |  |  |  | 0.2 ug/L |  |  | 80-125% |  | 80-124% |  | 0-40%
1,4-Dichlorobenzene |  |  |  |  | 0.2 ug/L |  |  | 80-125% |  | 80-120% |  | 0-40%
2-Butanone (MEK) |  |  |  |  | 4 ug/L |  |  | 49-150% |  | 80-150% |  | 0-40%
2-Hexanone |  |  |  |  | 4 ug/L |  |  | 75-121% |  | 76-133% |  | 0-40%
4-Methyl-2-Pentanone |  |  |  |  | 4 ug/L |  |  | 63-133% |  | 67-141% |  | 0-40%
Acetone |  |  |  |  | 4 ug/L |  |  | 20-150% |  | 21-150% |  | 0-40%
Acrylonitrile |  |  |  |  | 0.07 ug/L |  |  | 62-150% |  | 80-150% |  | 0-40%
Benzene |  |  |  |  | 0.2 ug/L |  |  | 80-133% |  | 80-129% |  | 0-40%
Bromochloromethane |  |  |  |  | 0.2 ug/L |  |  | 80-134% |  | 80-136% |  | 0-40%
Bromodichloromethane |  |  |  |  | 0.2 ug/L |  |  | 80-136% |  | 80-136% |  | 0-40%
Bromoform |  |  |  |  | 0.2 ug/L |  |  | 66-132% |  | 62-138% |  | 0-40%
Carbon Disulfide |  |  |  |  | 0.2 ug/L |  |  | 80-139% |  | 80-131% |  | 0-40%
Carbon Tetrachloride |  |  |  |  | 0.2 ug/L |  |  | 80-140% |  | 80-134% |  | 0-40%
Chlorobenzene |  |  |  |  | 0.2 ug/L |  |  | 80-126% |  | 80-120% |  | 0-40%
Chloroethane |  |  |  |  | 0.2 ug/L |  |  | 80-128% |  | 80-126% |  | 0-40%
Chloroform (1,1,1-Trichloromethane) |  |  |  |  | 0.2 ug/L |  |  | 80-129% |  | 80-128% |  | 0-40%
Chloromethane |  |  |  |  | 0.2 ug/L | 0.416 |  | 75-150% | 66/62 | 75-150% |  | 0-40%
cis-1,2-Dichloroethene |  |  |  |  | 0.2 ug/L | Y |  | 80-133% | Y | 80-131% |  | 0-40%
cis-1,3-Dichloropropene |  |  |  |  | 0.2 ug/L |  |  | 80-119% |  | 80-121% |  | 0-40%
Dibromochloromethane |  |  |  |  | 0.2 ug/L |  |  | 80-127% |  | 80-129% |  | 0-40%
Dibromomethane (Methylene bromide) |  |  |  |  | 0.2 ug/L |  |  | 80-128% |  | 80-136% |  | 0-40%
Dichlorodifluoromethane |  |  |  |  | 0.2 ug/L |  |  | 57-142% |  | 46-129% |  | 0-40%
Ethylbenzene |  |  |  |  | 0.2 ug/L |  |  | 80-141% |  | 80-133% |  | 0-40%
m,p-Xylene |  |  |  |  | 0.2 ug/L |  |  | 80-135% |  | 80-135% |  | 0-40%
Methyl Bromide (Bromomethane) |  |  |  |  | 0.2 ug/L |  | 68 | 70-150% |  | 63-146% |  | 0-40%
Methyl Iodide (Iodomethane) |  |  |  |  | 0.2 ug/L |  | Y | 20-150% |  | 20-150% |  | 0-40%
Methylene Chloride |  |  |  |  | 0.2 ug/L |  |  | 80-126% |  | 80-126% |  | 0-40%
o-Xylene |  |  |  |  | 0.2 ug/L |  |  | 80-139% |  | 80-134% |  | 0-40%
Styrene |  |  |  |  | 0.2 ug/L |  |  | 80-134% |  | 66-148% |  | 0-40%
Tetrachloroethylene |  |  |  |  | 0.2 ug/L |  |  | 80-125% |  | 80-119% |  | 0-40%
Toluene |  |  |  |  | 0.2 ug/L |  |  | 80-134% |  | 80-127% |  | 0-40%
trans-1,2-Dichloroethene |  |  |  |  | 0.2 ug/L |  |  | 80-137% |  | 80-133% |  | 0-40%
trans-1,3-Dichloropropene |  |  |  |  | 0.2 ug/L |  |  | 80-119% |  | 80-121% |  | 0-40%
trans-1,4-Dichloro-2-butene |  |  |  |  | 100 ug/L |  |  | 58-122% |  | 64-123% |  | 0-40%
Trichloroethylene (1,1,2-Trichloroethylene |  |  |  |  | 0.2 ug/L |  |  | 80-133% |  | 80-131% |  | 0-40%
Trichlorofluoromethane |  |  |  |  | 0.2 ug/L |  |  | 72-150% |  | 63-144% |  | 0-40%
Vinyl Acetate |  |  |  |  | 0.2 ug/L |  |  | 54-126% |  | 45-135% |  | 0-40%
Vinyl Chloride |  |  |  |  | 0.02 ug/L |  |  | 73-128% |  | 73-128% |  | 0-40%
4-Bromofluorobenzene | Surrogates |  |  |  |  | Y | 84-115%
d4-1,2-Dichloroethane | [5 empty] | Y | 80-120%
d8-Toluene | [5 empty] | Y | 88-111%
VOA Trip Blank Detects:
Associated Samples (to VTRPs Detects):
Data Anomalies Noted in 2-Yr Comparison:
Deviations from QAPP:
Corrective Action Advised:
Quality Control Specs: | Environmental Monitoring SAP for Cedar Hills Regional Landfill Appendix A??
1 Contract Specific Reporting Limits | [16 empty] | NA - Not Analyzed
Note: All storage, preservation and hold times were primairly based on info provided by KCEL in the "Sample_preservation_minimum_volumes._19-dec-14.xlsx", located in the S:\CLEE folder. All other info not provided in this document was researched using appropriate EPA Method Documents.

## Historical Data
chemical_name | chemical_name | chemical_name | 1,1,1,2-TETRACHLOROETHANE | 1,1,1-TRICHLOROETHANE | 1,1,2,2-TETRACHLOROETHANE | 1,1,2-TRICHLOROETHANE | 1,1-DICHLOROETHANE | 1,1-DICHLOROETHENE | 1,2,3-TRICHLOROPROPANE | 1,2-DIBROMO-3-CHLOROPROPANE | 1,2-DIBROMOETHANE | 1,2-DICHLOROBENZENE | 1,2-DICHLOROETHANE | 1,2-DICHLOROPROPANE | 1,4-DICHLOROBENZENE | 2-BUTANONE | 2-HEXANONE | 4-METHYL-2-PENTANONE | ACETONE | ACRYLONITRILE | Alkalinity, Bicarbonate | ALKALINITY, TOTAL (AS CACO3) | AMMONIA AS N | ANTIMONY | ARSENIC | BARIUM | BENZENE | BERYLLIUM | BROMOCHLOROMETHANE | BROMODICHLOROMETHANE | BROMOFORM | BROMOMETHANE | CADMIUM | CALCIUM | CARBON DISULFIDE | CARBON TETRACHLORIDE | CHLORIDE | CHLOROBENZENE | CHLORODIBROMOMETHANE | CHLOROETHANE | CHLOROFORM | CHLOROMETHANE | CHROMIUM | CIS-1,2-DICHLOROETHENE | CIS-1,3-DICHLOROPROPENE | COBALT | COPPER | DIBROMOMETHANE | DICHLORODIFLUOROMETHANE | ETHYLBENZENE | IRON | LEAD | M & P XYLENE | MAGNESIUM | MANGANESE | MERCURY | METHYL IODIDE | METHYLENE CHLORIDE | NICKEL | NITRATE | O-XYLENE | POTASSIUM | SELENIUM | SILVER | SODIUM | STYRENE | SULFATE | TETRACHLOROETHENE | THALLIUM | TOLUENE | TOTAL DISSOLVED SOLIDS | TOTAL ORGANIC CARBON | TOTAL SUSPENDED SOLIDS | TRANS-1,2-DICHLOROETHENE | TRANS-1,3-DICHLOROPROPENE | TRANS-1,4-DICHLORO-2-BUTENE | TRICHLOROETHENE | TRICHLOROFLUOROMETHANE | VANADIUM | VINYL ACETATE | VINYL CHLORIDE | ZINC
report_result_unit | report_result_unit | report_result_unit | ug/l | ug/l | ug/l | ug/l | ug/l | ug/l | ug/l | ug/l | ug/l | ug/l | ug/l | ug/l | ug/l | ug/l | ug/l | ug/l | ug/l | ug/l | mg/l | mg/l | mg/l | mg/l | mg/l | mg/l | ug/l | mg/l | ug/l | ug/l | ug/l | ug/l | mg/l | mg/l | ug/l | ug/l | mg/l | ug/l | ug/l | ug/l | ug/l | ug/l | mg/l | ug/l | ug/l | mg/l | mg/l | ug/l | ug/l | ug/l | mg/l | mg/l | ug/l | mg/l | mg/l | mg/l | ug/l | ug/l | mg/l | mg/l | ug/l | mg/l | mg/l | mg/l | mg/l | ug/l | mg/l | ug/l | mg/l | ug/l | mg/l | mg/l | mg/l | ug/l | ug/l | ug/l | ug/l | ug/l | mg/l | ug/l | ug/l | mg/l
fraction | fraction | fraction | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | T | T | T | N | T | N | N | N | N | T | T | N | N | N | N | N | N | N | N | T | N | N | T | T | N | N | N | T | T | N | T | T | T | N | N | T | N | N | T | T | T | T | N | N | N | T | N | N | N | N | N | N | N | N | N | T | N | N | T
sys_loc_code | sys_sample_code | sample_date
FS-3 | LFS3150616- | 2015-06-16 17:05 | 0.2 U | 0.2 U | 0.2 U | 0.2 U | 0.2 U | 0.2 U | 0.2 U | 1 U | 0.2 U | 0.2 U | 0.2 U | 0.2 U | 0.38 T | 4 U | 4 U | 4 U | 4.8 T | 0.07 U | 138 | 138 | 77.8 | 0.02 T | 0.001 U | 0.0583 | 0.2 U | 0.001 U | 0.2 U | 0.2 U | 0.2 U | 0.2 U | 0.002 U | 33.3 | 0.2 U | 0.2 U | 245 | 0.2 U | 0.2 U | 0.2 U | 0.2 U | 0.21 BT | 0.0068 T | 0.23 T | 0.2 U | 0.0074 T | 0.012 T | 0.2 U | 0.2 U | 0.2 U | 1.39 | 0.001 U | 0.2 U | 15.7 | 0.267 | 0.0001 U | 0.2 U | 0.2 U | 0.0282 | 18.7 | 0.2 U | 47.7 | 0.001 U | 0.003 U | 181 | 0.2 U | 344 | 0.2 U | 0.001 U | 0.2 U | 1060 | 66.4 | 2.6 | 0.2 U | 0.2 U | 100 U | 0.2 U | 0.2 U | 0.002 U | 0.2 U | 0.02 U | 0.0757

## Field Data
param_code | param_code | param_code | COND (F) | DO | ORP | PH (F) | TEMP
param_unit | param_unit | param_unit | uS/cm | mg/l | mV | pH units | deg c
sys_loc_code | sys_sample_code | measurement_date
FS-3 | LFS3150616- | 2015-06-16 17:05 | 1977 | 4.79 | 75.3 | 6.94 | 13.5

## Hold Time
Sample ID | Analysis Method (Copied from Tst.csv) | Prep Method (Copied from Tst.csv) | Analyte | Date/Time Collected | Date/Time Prepped | Date/Time Analyzed | Analysis Hold Time (Hours) | Analysis Hold Time (Days) | Maximum Allowable Analysis Hold Time (Days) | Analysis Hold Time vs. Max Allowable Analysis Hold Time (Days) | Prep Hold Time (Days) | Maximum Allowable Prep Hold Time (Days)
LFS3150616- | SW6020A | E200.8 | Total Metals | 2015-06-16 17:05 | 2015-06-29 10:00 | 2015-06-30 12:28 | 331.3833 | 13 | 180 | 13 | 12 | NA
LFS3150616- | SW7470A | E245.1 | Total Mercury | 2015-06-16 17:05 | 2015-06-24 11:30 | 2015-06-25 15:20 | 214.25 | 8 | 180 | 8 | 7 | NA
LFS3150616- | KEROUEL AMINOT 1997 | KEROUEL AMINOT 1997 | Ammonia-Nitrogen | 2015-06-16 17:05 | 2015-06-17 13:00 | 2015-06-19 10:37 | 65.5333 | 2 | 28 | 2 | 0 | NA
LFS3150616- | SM2320-B | SM2320-B | Alkalinity | 2015-06-16 17:05 | 2015-06-22 16:30 | 2015-06-22 16:30 | 143.4167 | 5 | 14 | 5 | 5 | NA
LFS3150616- | SM2320-Bi | SM2320-B | Alkalinity-Bi | 2015-06-16 17:05 | 2015-06-22 16:30 | 2015-06-22 16:30 | 143.4167 | 5 | 14 | 5 | 5 | NA
LFS3150616- | SM2540-C | SM2540-C | TDS | 2015-06-16 17:05 | 2015-06-22 13:35 | 2015-06-24 09:11 | 184.1 | 7 | 7 | 5 | 5 | NA
LFS3150616- | SM2540-D | SM2540-D | TSS | 2015-06-16 17:05 | 2015-06-23 12:52 | 2015-06-23 15:17 | 166.2 | 6 | 7 | 6 | 6 | NA
LFS3150616- | SM4110B CL | SM4110B CL | Chloride | 2015-06-16 17:05 | 2015-06-23 12:45 | 2015-06-23 12:45 | 163.6667 | 6 | 28 | 6 | 6 | NA
LFS3150616- | SM4110B SO4 | SM4110B SO4 | Sulfate | 2015-06-16 17:05 | 2015-06-23 12:45 | 2015-06-23 12:45 | 163.6667 | 6 | 28 | 6 | 6 | NA
LFS3150616- | SM4500-NO3F NO3 | SM4500-NO3F NO3 | Nitrite + Nitrate Nitrogen | 2015-06-16 17:05 | 2015-06-17 13:00 | 2015-06-17 14:33 | 21.4667 | 0 | 2 | 0 | 0 | 180
LFS3150616- | SM5310-B | SM5310-B | TOC | 2015-06-16 17:05 | 2015-06-25 14:04 | 2015-06-25 14:04 | 212.9833 | 8 | 28 | 8 | 8 | 28
LFS3150616- | SW8260C | SW5030B | VOA's | 2015-06-16 17:05 | 2015-06-18 15:31 | 2015-06-18 15:31 | 46.4333 | 1 | 14 | 1 | 1 | 14
LFS3150616- | SW846 8260C TIC | SW5030B | TIC's | 2015-06-16 17:05 | 2015-06-18 15:31 | 2015-06-18 15:31 | 46.4333 | 1 | 14 | 1 | 1 | 14
[9 empty] | Green = | Within Hold Time
[9 empty] | Red = | Outside Hold Time